## Ongoing
# | Task | Category | Status | SW Version | Date | Additional Comments | Jerome's Comment
1 | Results repeatability issue when running two consecutive power sweeps. | Bug | Not started |  |  |  | Only occurs when running in CW with no power calibration.
2 | In the CSV File output, fix the wrong display for pulse CW (Always says "NO"), and add two lines underneath showing pulse width and pulse period. | Bug / Improvement | Not started
3 | Increase the number of points in power interpolation plots | Bug | Done | 1.5.5 |  | vs Pin, Pout plots have 10x points | It works, but, just for you to know : It does require to delete and plot again if you change the number of points in the power sweep. I have the feeling this issue is related to issue number 6. Can you confirm? Nick: Yes, It is related to item #6. This has been fixed as well. Jerome : Works, but now it created issue 1 above
4 | Resolve the x-axis scaling issue caused by a change in measurement conditions | Bug | Done | 1.5.6 |  | Discussed this with Mike Leffel. Was given advice on how to force the rescale via SCPI commands. Both the number of points and x-axis limits are updated correctly now (see Jerome's comment for item 2 | Let us know about Munich response to this. Jerome : Works, but now it created issue 1 above
5 | "Autoscale All" functionality | Improvement | Not started (Done?) | ??? | ??? | Could still discuss making a separate app for this. Maybe put it on hold for now? | Interestingly enough, it seems to be resolved in current version. I feel it got resolved along with a previous issue. Jerome : Works, but I want to make sure we eliminated issue 1 before we consider it closed.
6 | End of sweep conditions (RF Off with power displayed at the bottom of plots) | Bug | In Progress* | 1.6 | 2016-02-05 | This depends on the development team in Munich. I can make the request, but I don't know if/when it would get implemented. A work-around Greg recommended is to display the power range in the diagram title. Not sure if that would work since in general the ranges in the diagram may not all be the same. Another idea is to use markers at the trace endpoints. | I think it would be good to feed Munich with this idea. One thing I managed to achieve here is to switch RF Off and then switch it back on. The RF power doesn't seem to come back on unless to trigger another sweep, but the power is displayed properly below each plots. Nick: As long as you use the RF off between sweeps option that I relayed to you before, then yes, this should be a useful work-around. Jerome : Works well. Is this the final version? If so, we can say this item is closed.
7 | Normalize Phase of Pout to lowest power measurement level and call this quantity AMPM. Make this available for plot in the GUI just like Pin, Pout, Sxy (on the Y-axis only, there is no point in making it available on the X-axis). | Improvement | In Progress | 1.7 | 2016-02-10
8 | Report AMPM in the CSV File | Improvement | In Progress | 1.7
9 | Message when plot of value exceeding the measurement boundaries | Improvement | Not started | 1.7 |  | The GUI will not let you input incorrect values: any at value that is not possible is rounded to the nearest possible value. If you watch the value as you enter it, you should see this. That said, I will add a warning in the next version. | Sounds good
10 | Minimized GUI ** | Improvement | In progress | 2 | 2016-02-17 | On hold for critical bug fixes. | See example figure below. We've been thinking of what it should look like. All the functionnalities should be accessible from there. For instance, the Plot window should remain accessible at all time and not only after a new measurement is captured. - Measure : Triggers the measurement - Meas. Config. - Plot - Export : Grant access to each of the specified windows.
11 | Control the values entered in the GCA setup, mainly for the frequency sweep, but possibly for the power sweep also. If the frequencies we enter don't exist, or if the sweep falls on values that don't exist, the software crashes (the software doesn't interpolate between frequencies). | Bug | Not started | 2 |  | Need clarification | Jerome : Typical crash case : VNA cal from 2 to 3 GHz, 11 frequency points. If I request GCA sweep from 2 to 3 GHz, 8 frequency points, it crashes (but 3-6 or 11 points will work).
12 | Speed optimization | Improvement | In progress | 2.1 | 2016-02-24 | On hold for critical bug fixes.
13 | Write a user manual for the application | Deliverable | Not started | 2.1
**

## Completed
# | Task | Category | Status | Proposed Date for Resolution / SW Version | Additional Comments | Jerome's Comment
| Thorough testing of permutations in table below* | Validation / Bug | Done* | 1.5.5* | Finished testing the various trace combinations, but I couldn't reproduce the intermittent error when plotting vs Pout. I did fix an issue you saw when moving from one X-axis type to another (say Pin to Frequency, and vice versa). Unfortunately the fix for now is for me to delete then recreate the trace entirely. I expect these trace changes will only need to be done during the initial configuration, so hopefully this isn't a big issue. | I couldn't reproduce it as well. We will have to keep an eye on this. Deleting and re-creating the trace isn't really an issue here as the change in trace format is significant and potentially requires rescaling.
| Move plots to a different figure | Bug | Done | 1.5.5 | Unfortunately, I had to reenable the data traces for the firmware to allow you to move the memory trace correctly | Works indeed
| Allow plot of gain and power values on the same figure | Bug | Done | 1.5.5 | You may have run into either the problem above, or the limitation of at most 20 traces per diagram (which includes data traces). I was able to do this. | Works indeed. We won't try to stack 20 plots on top of each others.
| Change plot nomenclature to something clearer (ie S11, S21, S12, S22) | Improvement | Done | 1.5.5 |  | Looks good
| CSV file add phase of Pout (same as phase of S21) | Improvement | Done | 1.5.5 | FWIW, also added phase of Pin (0 deg) | Looks good
| CSV file comma-delimeter | Improvement | In Progress | 1.6
| CSV file add settings as header (combine previous files into CSV) | Improvement | In Progress | 1.6
| CSV file add pulse definition info | Improvement | In Progress | 1.6
*