## Slide 1: PA Compression Test User Guide
Nick Lalic
VNA Software Developer
Rohde & Schwarz North America
March 30th, 2016

## Slide 2: Physical Setup
Example physical setup
Actual setup may vary with measurement needs (power level, pulsed RF, etc)
DUT S(f, Pin)
S1
a1
b1
Port1
S2
a2
b2
Port2
Driver

## Slide 3: Calibration
For full two-port S-Parameters:
  UOSM
  TOSM
  Any other full calibration
S-Parameter Calibration
DUT S(f, Pin)
S1
a1
b1
Calibration Planes
Port1
S2
a2
b2
Port2
Driver

## Slide 4: Calibration
For S11, S21 only:
  One Path Two Port calibration
S-Parameter Calibration
DUT S(f, Pin)
S1
a1
b1
Calibration Planes
Port1
S2
a2
b2
Port2
Driver

## Slide 5: Calibration
Standard Power cal of Port 1 source (S1), reference receiver (a1)
Measure Pin (a1) along with S-Parameters
Calculate compression point, Pout from Pin and S21
No need to power calibrate Port 1 measurement receiver (b1) or Port 2
Power Calibration
Pin
Pout = Pin + S21
DUT S(f, Pin)
S1
a1
b1
Power Cal. Plane
Port1
S2
a2
b2
Port2
Driver

## Slide 6: Measurement
Because the measured power (PIn, a1) is being used, the measurement points vs power do not have to (and in general will not) perfectly line up on a grid.
Compression is calculated using measured S-Parameters, PIn
Note: Because of limitations of the VNA, traces may be re-interpolated onto a linearly spaced x-axis as necessary. Raw data (Export) is not effected.
Pin
Gain

## Slide 7: Graphical User Interface
Settings tab
Export results to file (Touchstone, CSV)
Toggle Mini GUI mode
Exit application
Start Measurement
Frequency settings
Power settings
Miscellaneous settings

## Slide 8: Graphical User Interface
Traces tab
Export results to file (Touchstone, CSV)
Toggle
Mini GUI
mode
Exit
Write traces
to VNA
Trace list
Add/Remove
Trace

## Slide 9: Mini Graphical User Interface
Measure and plot traces
Export results (Touchstone, CSV)
Exit
Toggle Mini-GUI

## Slide 10: Settings
Start Frequency
Stop Frequency
Points
IF Bandwidth
Frequency settings are strictly obeyed
Frequency Settings

## Slide 11: Settings
Start Power
Stop Power
Points
Compression level (typically 1-3 dB)
Stop at Compression: Stops increasing measurement power level at or just after compression is achieved. DUT is driven at most ΔP over compression
Note: Since the compression point of most PAs varies vs frequency, Stop at compression will result in a different number of power measurement points per each frequency (ie the resulting matrix will not be “square”)
Power Settings

## Slide 12: Settings
Include Gain Expansion: When unchecked, the gain at the lowest power level is used for each frequency. For PAs that include gain expansion (gain can increase as power increases), this option finds the power level of maximum gain, and uses that gain value to find compression.
Power Settings

## Slide 13: Settings
Post-Condition
  RF Off: Stops all sweeps and turns power off between sweeps after measurements end
  None: Does nothing after measurements end
Power Settings

## Slide 14: Settings
Note: The RF Off post condition uses a setting called “Reduce RF Power between sweeps”. This setting persists through instrument preset. To put the instrument back in it’s factory default state (power on at last measurement point between sweeps), the setting can be found at: System Config (hard key) -> System Config (soft key) -> Power (tab) -> Uncheck “Power Reduction at Sweep End”
Power Settings

## Slide 15: Settings
(Reference) Channel:
  Settings outside the scope of PA Compression Test should be set in a reference channel before measurement
  These settings are used for the measurement
Settings taken from reference channel:
  User defined ports
  Pulsed RF power
  Calibration
  Power calibration
  … Anything not explicitly overwritten in Settings
Miscellaneous Settings

## Slide 16: Settings
Output Port, Input Port
Sweep Type:
  Frequency
    A segmented frequency sweep is performed for a fixed power level. The power level is increased with each sweep.
    When Stop at Compression is selected, frequencies that have reached compression are removed from subsequent segmented frequency sweeps.
  Power
    For a fixed frequency, a power sweep is performed per Power Settings
Miscellaneous Settings

## Slide 17: Traces
Y:
  S11, S21, S12, S22
  AM-PM Conversion
  Pin
  Pout
X:
  Frequency
  Pin
  Pout
At:
  Compression
  Max. Gain
  Frequency value
  Power value
Trace Options

## Slide 18: Traces
Initially, traces are created in new diagram(s) on the VNA
User then configures each trace (diagram, format, markers etc)
Subsequent replotting will update trace values without disturbing customization.

## Slide 19: Traces
Note: For the purposes of this application, AMPM is defined as: AMPM = unwrapped_phase(S21) – unwrapped_phase(S21[PLow])
AMPM can be plotted versus various x-axis
Example: AMPM vs POut @ 500 MHz, 3 GHz and 6 GHz

## Slide 20: Export
Results are exported into a ZIP file
Results are exported in two formats:
  Touchstone 1.0 files – One per power level
  Compression CSV – Summary of compression point results
  Raw Data CSV file – containing all raw measurement points

## Slide 21: Export
Header
  VNA info
  Settings (Frequency, Power, Miscellaneous)
  RF Pulse info (if applicable)
Data Columns:
  Frequency
  Pin (dBm, deg)
  Pout (dBm, deg)
  AMPM (deg)
  S11, S12, S21, S22 (dB, deg)
CSV Format

## Slide 22: Contact
Please contact me with questions or comments
      Nick Lalic VNA Software Developer
      San Francisco, CA
      +1 (424) 200-2846 nick.lalic@rsa.rohde-schwarz.com
      http://vna.rs-us.net