## Slide 1: PA Compression Test User Guide
Nick Lalic
VNA Software Developer
Rohde & Schwarz North America
April 4th, 2016

## Slide 2: Sections
Hardware Configuration
Calibration
  ZVAX Pulse modulator on
  Pulse Gen: Constant High
  Trigger: Free Run
  One Path Two Port
  Source Power Cal
Pulsed RF
  Pulse Type: Single Pulse
  Trigger: Rising Edge Pulse
  Trigger measurement: Point

## Slide 3: Hardware Configuration
DUT S(f, Pin)
S1
a1
b1
Calibration Planes
Port1
S2
a2
b2
Port2
Driver
Power Cal. Plane
ZVAX

## Slide 4: Hardware Configuration
ZVA USB -> ZVAX “USB from NWA”
ZVA Cascade -> ZVAX Cascade In
Rear Connections

## Slide 5: Hardware Configuration
RF Connections
VNA Port 1 Source Out to ZVAX Port 1 In

## Slide 6: Hardware Configuration
RF Connections
VNA Port 1 Source Out -> ZVAX Port 1 In
VNA Port 1 Ref In, Meas In from couplers
VNA Port 2 Meas In only (or normal port access)
DUT S(f, Pin)
Driver

## Slide 7: Calibration
We will leave the pulse modulator switch in the on position for the cal
Verify that the Src 1 Pulse Modulator is in the ZVAX path
  Mode (hard key) -> More (if/as necessary) -> ZVAX Path Config
Setting up

## Slide 8: Calibration
You should hear a click (physical switches)
ZVAX Port 1 Pulse Mod should be on
Setting up

## Slide 9: Calibration
Mode (hard key) -> Pulse Generator (on, checked)
Setting up

## Slide 10: Calibration
Click Def Pulse Generator…
Setting up

## Slide 11: Calibration
Pulse Type to Constant High
Setting up

## Slide 12: Calibration
Sweep (hard key) -> Trigger -> Free Run
At this point the analyzer should be sweeping freely, without pulsing
Setting up

## Slide 13: Calibration
Cal (hard key) -> Start Cal -> Two-Port P1 P2 -> One Path Two-Port…
Follow on-screen directions
One Path Two Port
DUT S(f, Pin)
S1
a1
b1
Calibration Planes
Port1
S2
a2
b2
Port2
Driver
ZVAX

## Slide 14: Calibration
Cal (hard key) -> Start Power Cal -> Source Power Cal…
Power Cal
DUT S(f, Pin)
S1
a1
b1
Port1
S2
a2
b2
Port2
Driver
Power Cal. Plane
ZVAX

## Slide 15: Calibration
Power Cal
Click Modify Settings…

## Slide 16: Calibration
Power Cal
Set Cal Offset to gain of driver amplifier
Set Maximum Number of Readings (iterations), Tolerance as appropriate
Include Reference Receiver Cal (Important)
Click OK

## Slide 17: Calibration
Power Cal
Auto Zero power sensor (optional)
Take Cal Sweep
  With iteration, source should converge on expected values
Close

## Slide 18: Pulsed RF
Mode (hard key) -> Def Pulse Generator
Set Pulse Type to Single Pulse
Set Pulse Width, Period

## Slide 19: Pulsed RF
Note: Pulse width must be longer than sampling time
Reference table:

## Slide 20: Pulsed RF
Sweep -> Trigger -> Pulse Gen…
Rising Edge Pulse
Pulse Gen... Should be selected

## Slide 21: Pulsed RF
Trigger Settings
Triggered Measurement Sequent to Point
Trigger Delay (Optional)

## Slide 22: Contact
      Nick Lalic VNA Software Developer
      San Francisco, CA
      +1 (424) 200-2846 nick.lalic@rsa.rohde-schwarz.com
      http://vna.rs-us.net
      Greg Bonaguide Senior Product Line Engineer
      Spectrum Analyzers
      Vector Network Analyzers
      Northeastern Region
      +1 (774) 2820028 greg.bonaguide@rsa.rohde-Schwarz.com
      Phone Support: +1 (888) TEST-RSA